## Slide 1: PXI Overview: Chassis
Chassis
  Special slot for controller or remote control card
  Special slot for star trigger (optional)
  The rest are peripheral slots
Power
Cooling
Isolated ground

## Slide 2: PXI Overview: Slot types
PXI was based on PCI
PXIe is based on PCIe
PXI-H hybrid slot supports both PXIe and PXI
PXIe did not obsolete PXI – both still used
PXIe is used for high-speed measurements
Physical size
  3U
  6U
  May take up multiple slots

## Slide 3: Minimum requirements to build a test
Voltage measurement
  We mean open circuit voltage typically
Resistance measurement
  Can tell you more about connection than open circuit voltage
Arbitrary voltage source
  Turn devices on and off
  Turn relays on
  Avoid using for continuity tests, prefer resistance tests
Current measurement (rare)
Being able to construct arbitrary circuits
Notes: First, those are minimum requirements to write tests. You want to measure voltage and resistance. Ideally, we would always measure resistance because it is more accurate, but it’s not always possible. Measuring an open circuit voltage on a pin, for example, tells you that it’s the correct potential, but does not tell you if it’s a high resistance connection (poor crimp, not seated terminal, etc…). Resistance test is more accurate, but it may not be possible because we don’t have access to both ends of the wire, or if there’s current going through the wire. Voltage measurement is used to test relays outputs, where as resistance measurement is used from grounds, and in places where we have access to both ends of the wire. There’s an additional requirement here: an arbitrary adjustable voltage source. This is important, because it can make it very easy to build tests if you could generate an arbitrary voltage on command. In circuit analysis, a test voltage source is an important measurement tool that can be used to characterize a circuit. For example, you may apply a test voltage source and measure current to find out impedance. Current measurement may be important for testing certain components, such as verifying sensor current sources, measuring output or input impedance, estimating power consumption. It’s important when you need it, but for Tiger may be optional. Last requirement, you want to be able to construct arbitrary test circuits in software. In this is the topic of the next series of slides.

## Slide 4: Matrix PXI Modules
Grid arrangement of relays
Each cross-point is a relay connecting row and column
Notes: To build arbitrary circuits we need a matrix. Matrix is a grid arrangement of relays. Typically still uses PCI(not PXIe)

## Slide 5: How to perform routing using a matrix?
Each matrix has a long X dimension and a shorter Y dimension
You may try to place test resources such as DMM and voltage sources on a short axis
You will find this has shortcomings!
Instead, just place all nodes on X axis
Much simpler and it will work…
X
Y
Notes: How do you use a matrix to build tests? When you look at different modules, the first thing you notice they have one short and one long dimension. You may try to put your scarce resources such as DMM probes or voltage sources on the short dimension and then put the device under test pins on the long dimension. This would work to a point, but soon you will realize you may want to duplicate some test resources onto the long axis to be able to build more types of circuits. Ultimately you can duplicate all your resources and have an 8x8 square dedicated to test resource pins. This has one problem though: the number of test resources must be less than the Y dimension of the matrix. Additionally, if you have more Y rows than the test resources you don’t have a way of using the extra rows, they are just dead wait.

## Slide 6: How to perform routing using a matrix?
The long axis is assigned to both Device Under Test (DUT) and test resource nodes.
Do not place anything on Y axis…
… until it is time to route!
X
Y

## Slide 7: Aside Q: What is a net? A: A net is a description of connectivity.
Formally, a net is a set of connected nodes.
Circuit may contain multiple nets.
Nets are disjoint sets, meaning a node can only belong to one net.

## Slide 8: How to perform routing using a matrix?
Use Y-axis rows to make arbitrary connections between X-axis columns
Each net requires a single row
Notes: The best way to route using a matrix is to keep track of nets. Each net requires a single row to implement it. For each net, select a row. Go over each node in the net and find the intersection of its column with the row. Turn the relay at intersection on to connect it. Each row acts as a set and by connecting columns to it you add it to the set.

## Slide 9: Example: Connect J2-32 to GND
Pick an unused row
Intersect J2-32 column with the row
Turn the relay at intersection on

## Slide 10: How big does the matrix need to be?
Tiger control box requires 204 targets
In addition to 7 tester targets
For a total of 211 targets on X-axis
Y-axis requirements depend on number of simultaneous nets in a test
Maximum Y-axis usage was 4 in simulations
(Actual maximum row usage was 5 with no optimizations)
A Y-axis of 6 would be preferred.

## Slide 11: Extending a matrix
Find a matrix module with desirable Y dimension
Buy as many modules as needed to extend the X dimension.
Connect the rows (Ys) of the modules into one large matrix.
Notes: This matrix is ending up very long. Where do you get a matrix with over 200 columns? You get one by connecting rows of smaller matrices.

## Slide 12: Build your own matrix router in two LISP calls:
Routing matrix was implemented in LISP in two functions
First one builds nets from connections
  A (connection a b) is a binary relation between two nodes

## Slide 13: Build your own matrix router in two LISP calls:
Second one, flips the relays on…

## Slide 14: Test resource nodes
The 7 test targets are dedicated to test resources
DMM_V_HI is Voltage Probe +
DMM_LO is DMM Ground
DMM_I_HI is DMM Current Probe +
PS_RAIL1,2,3 are adjustable power supply controlled by the test
GND

## Slide 15: Test resource nodes
DMM_V_HI
DMM_LO
DMM_I_HI

## Slide 16: Making a basic voltage measurement
Note on LISP syntax:
Calling a function
(<function name> arg1 arg2 …)
Symbol syntax
`<symbol> also ‘<symbol>
COST: 2 ROWS

## Slide 17: Making a basic resistance measurement
COST: 2 ROWS

## Slide 18: Applying a voltage signal
COST: 1 ROW

## Slide 19: Performing a current measurement (not used for Tiger)
COST: 3 ROWS

## Slide 20: Test language
connect two nodes a and b together
measure voltage between a and b nodes
measure resistance from a to b
apply-voltage v to node a
Connectivity | Power | Measurement
(connect a b) | (power-supply-set rail v) | (dmm-v)
(disconnect a b) | (power-supply-on rail) | (dmm-i)
(disconnect-all) | (power-supply-off rail) | (dmm-r)
 | (apply-voltage a v) | (voltage a b)
 | (remove-voltage a) | (resistance a b)
 | (apply-ground a) | 

## Slide 21: Test language
connect two nodes a and b together
Take dmm-voltage reading depending on where DMM_V_HI and DMM_V_LO are connected
Take dmm-resistance reading depending on where DMM_V_HI and DMM_V_LO are connected
Manually control power-supply nodes with …-set, on, off
Connectivity | Power | Measurement
(connect a b) | (power-supply-set rail v) | (dmm-v)
(disconnect a b) | (power-supply-on rail) | (dmm-i)
(disconnect-all) | (power-supply-off rail) | (dmm-r)
 | (apply-voltage a v) | (voltage a b)
 | (remove-voltage a) | (resistance a b)
 | (apply-ground a) | 

## Slide 22: Expect is way of asserting correct behavior
expect takes a message and Boolean expression
expect fails the test when the Boolean expression is false and returns an error message to the test system.

## Slide 23: Simplifying common patterns
We can define more shorthand calls to simplify common patterns

## Slide 24: Simplifying common patterns
We can define more shorthand calls to simplify common patterns

## Slide 25: Real test examples: KSI Relay
Applies 48V and GROUND to turn control box “on”.
Applies voltage to coil of relay K4
Reads voltages at various relay outputs
Expects read voltage to be 48V +/- 1V
Return error message if voltage is out of range

## Slide 26: Real test examples: Ground test
Connects control box ground
Power is OFF!
Measures resistance from pin to ground

## Slide 27: Real test examples: CAN Terminator
Connects control box ground
Power is OFF!
Measures resistance end to end

## Slide 28: Real test examples: Voltage test
Turn the control box ON
Measures voltage on one of 24V pins on master bus

## Slide 29: Minimum requirements to build a test
Voltage measurement
Resistance measurement
Arbitrary voltage source
Current measurement (rare)
Being able to construct arbitrary circuits

## Slide 30: Pickering catalog

## Slide 31: 4x64 Matrix Option from Pickering
Having 4 rows in Y-axis is very limiting
Need x4 modules for 256 targets on X axis
2A current

## Slide 32: 6x64 Matrix Option from Pickering
Having 6 rows is useful
Need x4 modules for 256 targets on X axis
2A current

## Slide 33: 6x64 Matrix Option from Marvin
Having 6 rows in Y-axis is nice
Still need x4 modules for 256 targets on X axis
Only 500 mA current

## Slide 34: Minimum requirements to build a test
Voltage measurement
Resistance measurement
Arbitrary voltage source
Current measurement (rare)
Being able to construct arbitrary circuits
Notes: First, those are minimum requirements to write tests. You want to measure voltage and resistance. Ideally, we would always measure resistance because it is more accurate, but it’s not always possible. Measuring an open ciruit voltage on a pin, for example, tells you that it’s the correct potential, but does not tell you if it’s a high resistance connection (poor crimp, not seated terminal, etc…). Resistance test is more accurate, but it may not be possible because we don’t have access to both ends of the wire, or if there’s current going through the wire. Voltage measurement is used to test relays outputs, where as resistance measurement is used from grounds, and in places where we have access to both ends of the wire. There’s an additional requirement here: an arbitrary adjustable voltage source. This is important, because it can make it very easy to build tests if you could generate an arbitrary voltage on command. In circuit analysis, a test voltage source is an important measurement tool that can be used to characterize a circuit. For example, you may apply a test voltage source and measure current to find out impedance. Current measurement may be important for testing certain components, such as verifying sensor current sources, measuring output or input impedance, estimating power consumption. It’s important when you need it, but for Tiger may be optional. Last requirement, you want to be able to construct arbitrary test circuits in software. In this is the topic of the next series of slides.

## Slide 35: PXI Power Supply
Possible candidate for adjustable power supply
500 mA, up to 32V
Ideal test signal source
Not great for switching large relays

## Slide 36: PXI Power Supply
2 Channel
2A, 48V
Great, but could use more channels

## Slide 37: PXI SMUs: Precision Sources
Higher-precision sources are called SMU (Source Measure Units)
Typically, low voltages/currents

## Slide 38: Minimum requirements to build a test
Voltage measurement
Resistance measurement
Arbitrary voltage source
Current measurement (rare)
Being able to construct arbitrary circuits
Notes: First, those are minimum requirements to write tests. You want to measure voltage and resistance. Ideally, we would always measure resistance because it is more accurate, but it’s not always possible. Measuring an open ciruit voltage on a pin, for example, tells you that it’s the correct potential, but does not tell you if it’s a high resistance connection (poor crimp, not seated terminal, etc…). Resistance test is more accurate, but it may not be possible because we don’t have access to both ends of the wire, or if there’s current going through the wire. Voltage measurement is used to test relays outputs, where as resistance measurement is used from grounds, and in places where we have access to both ends of the wire. There’s an additional requirement here: an arbitrary adjustable voltage source. This is important, because it can make it very easy to build tests if you could generate an arbitrary voltage on command. In circuit analysis, a test voltage source is an important measurement tool that can be used to characterize a circuit. For example, you may apply a test voltage source and measure current to find out impedance. Current measurement may be important for testing certain components, such as verifying sensor current sources, measuring output or input impedance, estimating power consumption. It’s important when you need it, but for Tiger may be optional. Last requirement, you want to be able to construct arbitrary test circuits in software. In this is the topic of the next series of slides.

## Slide 39: PXI DMM (Digital Multimeters): M1212A – High End
M1412A
7.5 Digit
Costs $8793
Notes: Keysight is known for making high-end oscilloscopes and other lab test and measurement equipment.

## Slide 40: PXI DMM (Digital Multimeters): Other options from Keysight

## Slide 41: PXI DMM (Digital Multimeters): NI PXI-4065
6.5 Digit
Costs $2085

## Slide 42: Bench Top DMM (Digital Multimeters): Other options from Keysight
6.5 Digit
Ethernet, USB
Rear probe jacks
Costs $1865

## Slide 43: Bench Top DMM (Digital Multimeters): And then there’s Fluke
5.5 Digit
Ethernet, USB Serial
Rear probe jacks
Costs $1360

## Slide 44: Summary
Rough wish list so far:
Matrix modules
  Four 4x64 modules
  6x64 preferred
  Need some with 2A capacity for power and large relays
  Signal can be 500mA capacity or less
Voltage source
  x1 3 rail test source supply 500mA general
  x1 high-power power supply 2A capability up to 48V for switching large relays and power input
DMM
Chassis
  PC-based if possible